## Sheet1
IMT
setting | Bpm | o2 | temp | flow
0.25 | 10 | 90.2 | 23.5 | 0
0.5 | 10 | 91.2 | 24.1 | 0
0.75 | 10 | 92.1 | 24.5 | 0.1
1 | 10 | 93 | 25 | 0.15
1.25 | 10 | 93.4 | 25.2 | 0.2
1.5 | 10 | 93.7 | 25.5 | 0.2
1.75 | 10 | 94 | 25.6 | 0.2
2 | 10 | 94.1 | 25.8 | 0.3
2.25 | 10 | 94.2 | 26 | 0.4
2.5 | 10 | 94.4 | 26.1 | 0.4
2.75 | 10 | 94.6 | 26.2 | 0.4
3 | 10 | 94.6 | 26.2 | 0.5
3.25 | 10 | 94.5 | 26.3 | 0.6
3.5 | 10 | 94.4 | 26.3 | 0.6
3.75 | 10 | 94.3 | 26.4 | 0.6
4 | 10 | 94 | 26.5 | 0.7
4.25 | 10 | 93.6 | 26.5 | 0.75
4.5 | 10 | 93.4 | 26.6 | 0.8
4.75 | 10 | 92.8 | 26.6 | 0.9
5 | 10 | 92.8 | 26.7 | 0.8

## TestBench Validation (2)
OLD zeolite cartridges
Last update: 13/07/2023 |  |  | CC-0 | [7 empty] | CC-20
|  |  | O2 (%) |  |  | FLOW (L/min) |  | Temperature (ºC) |  |  | O2 (%) |  |  | FLOW (L/min) |  | Temperature (ºC)
| Setting | RR (bpm) | P5 | IMT | TB | IMT | TB | Ambient | IMT | TB | P5 | IMT | TB | IMT | TB | Ambient | IMT | TB
| 5 | 25
| 4.5 | 25
| 3.75 | 25
| 2.5 | 25
| 1.25 | 25
| 0.5 | 25
| 5 | 10
| 4.5 | 10
| 3.75 | 10
| 2.5 | 10
| 1.25 | 10
| 0.5 | 10
Notes:
O2 must be superior to 87% in all coditions
Ref.
Selected manual measures taken on a P5

## TestBench Validation
OLD zeolite cartridges
Last update: 29/06/2023 |  |  | CC -0 | [9 empty] | CC -20 | [7 empty] | CC -10
|  |  |  | TESTBENCH |  |  | ProCheck | TESTBENCH | IMT | temperature |  |  |  | TEMPERATURE | ProCheck | TESTBENCH |  |  | IMT |  |  | ProCheck | IMT
| Setting | RR (bpm) | O2% (P5) | O2% | Min | Max | O2% | Flow L (ave) | Flow L (ave) | ambient | IMT | TESTBENCH | O2% (P5) | ambient | O2% | Temperature | O2% | Flow L (ave) | Temperature | Flow L (ave) | O2% (P5) | O2% | Flow L (ave)
q | 5 | 25 | 94 |  |  |  | 91.5 |  | 1.14 |  |  |  | 91.8 |  | 89 |  |  |  |  | 0.93 | 91.9 | 89 | 1
|  |  | 91.2 | 90.7 | 89.5735 | 91.8265 | 92.8 | 0.95 | 1.1 | 29 | 31.4 | 30.2 | 87.9 | 26.6 | 90.1 | 29.88 | 88.28 | 0.87 | 30.8 | 0.95
| 4.5 | 25 | 94.3 |  |  |  | 92 |  | 0.98 |  |  |  | 89.3 |  | 87 |  |  |  |  | 0.86 | 91 | 89 | 0.92
| 3.75 | 25 | 95 |  |  |  | 95.3 |  | 0.79 |  |  |  | 91.6 |  | 91.2 |  |  |  |  | 0.71 | 93.7 | 92 | 0.78
q | 2.5 | 25 | 98 |  |  |  | 96 |  | 0.55 |  |  |  | 97.1 |  | 96 |  |  |  |  | 0.42 | 95.8 | 95.7 | 0.45
| 1.25 | 25 | 97.6 |  |  |  | 96 |  | 0.27 |  |  |  | 97.5 |  | 96 |  |  |  |  | 0.21 | 97.9 | 96 | 0.24
q | 0.5 | 25 | 98 |  |  |  | 96 |  | 0.14 |  |  |  | 93.6 |  | 93.1 |  |  |  |  | 0.12 | 98 | 98.5 | 0.12
q | 5 | 10 | 96.5 |  |  |  | 95.5 |  | 1.05 |  |  |  | 91.7 |  | 89.7 |  |  |  |  | 0.99 | no | no | no
|  |  | 91.7 | 92.18 | 91.0535 | 93.3065 | 95.7 | 0.68 | 0.8 | 27.3 | 26.8 | 29.8 | [8 empty] | no | no | no
| 4.5 | 10 | 93.2 |  |  |  | 94.5 |  | 0.93 |  |  |  | 89.5 |  | 88.8 |  |  |  |  | 0.9 | no | no | no
[21 empty] | no | no | no
| 3.75 | 10 | 96.8 |  |  |  | 96 |  | 0.73 |  |  |  | 94.5 |  | 94.5 |  |  |  |  | 0.64 | no | no | no
[21 empty] | no | no | no
q | 2.5 | 10 | 98 |  |  |  | 98.4 |  | 0.45 |  |  |  | 96.3 |  | 96 |  |  |  |  | 0.39 | no | no | no
|  |  | 92.9 | 94.75 | 93.6235 | 95.8765 | 96 | 0.34 | 0.5 | 29.3 | 31.1 | 31.67 | [8 empty] | no | no | no
| 1.25 | 10 | 98 |  |  |  | 96 |  | 0.24 |  |  |  | 98 |  | 96 |  |  |  |  | 0.22 | no | no | no
[21 empty] | no | no | no
q | 0.5 | 10 | 98 |  |  |  | 96 |  | 0.12 |  |  |  | 98 |  | 98 |  |  |  |  | 0.12 | no | no | no
| 0.5 | 10 | 95.1 | 95.6 | 94.4735 | 96.7265 | 96 | 0.1 | 0.2 | 29.3 | 31.5 | 32.09
Notes:
O2 must be superior to 87% in all coditions
Ref.
Selected manual measures taken on a P5

## NEW
NEW zeolite cartridges
Last update: xx/06/2023 |  |  | CC -0 |  |  | CC -20 |  |  | CC -10
|  |  |  | ProCheck | IMT |  | ProCheck | IMT |  | ProCheck | IMT
| Setting | RR (bpm) | O2% (P5) | O2% | Flow L (ave) | O2% (P5) | O2% | Flow L (ave) | O2% (P5) | O2% | Flow L (ave)
| 5 | 25 | 94 | 91.5 | 1.14 | 91.8 | 89 | 0.93 | 91.9 | 89 | 1
| 4.5 | 25 | 94.3 | 92 | 0.98 | 89.3 | 87 | 0.86 | 91 | 89 | 0.92
| 3.75 | 25 | 95 | 95.3 | 0.79 | 91.6 | 91.2 | 0.71 | 93.7 | 92 | 0.78
| 2.5 | 25 | 98 | 96 | 0.55 | 97.1 | 96 | 0.42 | 95.8 | 95.7 | 0.45
| 1.25 | 25 | 97.6 | 96 | 0.27 | 97.5 | 96 | 0.21 | 97.9 | 96 | 0.24
| 0.5 | 25 | 98 | 96 | 0.14 | 93.6 | 93.1 | 0.12 | 98 | 98.5 | 0.12
| 5 | 10 | 96.5 | 95.5 | 1.05 | 91.7 | 89.7 | 0.99 | no | no | no
[9 empty] | no | no | no
| 4.5 | 10 | 93.2 | 94.5 | 0.93 | 89.5 | 88.8 | 0.9 | no | no | no
[9 empty] | no | no | no
| 3.75 | 10 | 96.8 | 96 | 0.73 | 94.5 | 94.5 | 0.64 | no | no | no
[9 empty] | no | no | no
| 2.5 | 10 | 98 | 98.4 | 0.45 | 96.3 | 96 | 0.39 | no | no | no
[9 empty] | no | no | no
| 1.25 | 10 | 98 | 96 | 0.24 | 98 | 96 | 0.22 | no | no | no
[9 empty] | no | no | no
| 0.5 | 10 | 98 | 96 | 0.12 | 98 | 98 | 0.12 | no | no | no
Notes:
O2 must be superior to 87% in all coditions
Ref.
Selected manual measures taken on a P5

## Clinical perf. (Christian info)
Clinical Perf. to be maintained (current P5) = Bolus par minute
Réglage de débit | Débit ( ml +/- 10%) |  |  | Concentration d'oxygène |  |  | 10 inspirations/minute |  |  | 20 inspirations/minute |  |  | 25 inspirations/minute
[7 empty] | ( +/- 3ml ou +/- 10%) |  |  | ( +/- 3ml ou +/- 10%) |  |  | ( +/- 3ml ou +/- 10%)
| Mini | Nom | Max | Mini | Nom | Max | min | nominal | max | min | nominal | max | min | nominal | max
0.25 | 47.25 | 52.5 | 57.75 | 87 | 90 | 96 | 4.7 | 5.25 | 5.7 | 2.4 | 2.6 | 2.8 | 1.9 | 2.1 | 2.3
0.5 | 94.5 | 105 | 115.5 | 87 | 90 | 96 | 9.4 | 10.5 | 13.6 | 4.7 | 5.2 | 5.7 | 3.8 | 4.2 | 4.6
0.75 | 141.8 | 158 | 173.25 | 87 | 90 | 96 | 14.1 | 15.7 | 17.3 | 7.1 | 7.9 | 8.7 | 5.7 | 6.3 | 6.9
1 | 189 | 210 | 231 | 87 | 90 | 96 | 18.9 | 21 | 23.1 | 9.5 | 10.5 | 13.6 | 7.6 | 8.4 | 9.2
1.25 | 236.3 | 263 | 288.75 | 87 | 90 | 96 | 23.5 | 26.2 | 28.8 | 11.7 | 13.1 | 14.4 | 9.4 | 10.5 | 11.5
1.5 | 283.5 | 315 | 346.5 | 87 | 90 | 96 | 28.3 | 31.5 | 34.6 | 14.1 | 15.7 | 17.3 | 11.3 | 12.6 | 13.8
1.75 | 330.8 | 368 | 404.25 | 87 | 90 | 96 | 33 | 36.7 | 40.3 | 16.5 | 18.3 | 20.1 | 13.2 | 14.7 | 16.1
2 | 378 | 420 | 462 | 87 | 90 | 96 | 37.8 | 42 | 46.2 | 18.9 | 21 | 27.2 | 15.1 | 16.8 | 18.5
2.25 | 425.3 | 473 | 519.75 | 87 | 90 | 96 | 42.4 | 47.2 | 51.9 | 21.2 | 23.6 | 25.9 | 17 | 18.9 | 20.8
2.5 | 472.5 | 525 | 577.5 | 87 | 90 | 96 | 47.25 | 52.5 | 57.7 | 23.6 | 26.2 | 28.8 | 18.9 | 21 | 23.1
2.75 | 519.8 | 578 | 635.25 | 87 | 90 | 96 | 51.9 | 57.7 | 63.4 | 25.9 | 28.8 | 31.7 | 20.8 | 23.1 | 25.4
3 | 567 | 630 | 693 | 87 | 90 | 96 | 56.7 | 63 | 69.3 | 28.4 | 31.5 | 40.8 | 22.7 | 25.2 | 27.7
3.25 | 614.3 | 683 | 750.75 | 87 | 90 | 96 | 61.3 | 68.2 | 75 | 30.6 | 34.1 | 37.5 | 24.5 | 27.3 | 30
3.5 | 661.5 | 735 | 808.5 | 87 | 90 | 96 | 66.1 | 73.5 | 80.8 | 33 | 36.7 | 40.4 | 26.4 | 29.4 | 32.3
3.75 | 708.8 | 788 | 866.25 | 87 | 90 | 96 | 70.8 | 78.7 | 86.5 | 35.4 | 39.3 | 43.2 | 28.3 | 31.5 | 34.6
4 | 756 | 840 | 924 | 87 | 90 | 96 | 75.6 | 84 | 92.4 | 37.8 | 42 | 54.4 | 30.2 | 33.6 | 37
4.25 | 803.3 | 893 | 981.75 | 87 | 90 | 96 | 80.2 | 89.2 | 98.1 | 40.1 | 44.6 | 49 | 32.1 | 35.7 | 39.2
4.5 | 850.5 | 945 | 1039.5 | 87 | 90 | 96 | 85 | 94.5 | 103.9 | 42.5 | 47.2 | 51.9 | 34 | 37.8 | 41.5
4.75 | 897.8 | 998 | 1097.3 | 87 | 90 | 96 | 89.7 | 99.7 | 109.6 | 44.8 | 49.8 | 54.8 | 35.9 | 39.9 | 43.9
5 | 945 | 1050 | 1155 | 87 | 90 | 96 | 94.5 | 105 | 115.5 | 47.3 | 52.5 | 67.9 | 37.8 | 42 | 46.2